|  | Document # | Date Effective |
| --- | --- | --- |
|  | LAT-TD-00844-D1 | Draft 07/18/02 |
|  | Prepared by(s) | Supersedes |
|  | Alex Moiseev | None |
| GLAST LAT Subsystem Technical Document |  |  |
|  | Subsystem/Office |
|  | Anticoincidence Detector Subsystem |
| Document Title |
| ACD Gain Calibration Test with Cosmic Ray Muons |

Gamma-ray Large Area Space Telescope (GLAST)

Large Area Telescope (LAT)

# ACD Gain Calibration Test with Cosmic Ray Muons

DRAFT

CHANGE HISTORY LOG

| Revision | Effective Date | Description of Changes |
| --- | --- | --- |

# Purpose

This study reports on test methods for verifying the performance of the ACD after assembly.

- Definitions and Acronyms

| ACD | The LAT Anti-Coincidence Detector Subsystem |
| --- | --- |
| ADC | Analog-to-Digital Converter |
| AEM | ACD Electronics Module |
| ASIC | Application Specific Integrated Circuits |
| BEA | Base Electronics Assembly |
| CAL | The LAT Calorimeter Subsystem |
| DAQ | Data Acquisition |
| EGSE | Electrical Ground Support Equipment |
| EMC | Electromagnetic Compatibility |
| EMI | Electromagnetic Interference |
| ESD | Electrostatic Discharge |
| FM | Flight Module |
| FMEA | Failure Mode Effect Analysis |
| FREE | Front End Electronics |
| GAFE | GLAST ACD Front End – Analog ASIC |
| GARC | GLAST ACD Readout Controller – Digital ASIC |
| GEVS | General Environmental Verification Specification |
| GLAST | Gamma-ray Large Area Space Telescope |
| HVBS | High Voltage Bias Supply |
| ICD | Interface Control Document |
| IDT | Instrument Development Team |
| I&T | Integration and Test |
| IRD | Interface Requirements Document |
| JSC | Johnson Space Center |
| LAT | Large Area Telescope |
| MGSE | Mechanical Ground Support Equipment |
| MLI | Multi-Layer Insulation |
| MPLS | Multi-purpose Lift Sling |
| PCB | Printed Circuit Board |
| PDR | Preliminary Design Review |
| PMT | Photomultiplier Tube |
| PVM | Performance Verification Matrix |
| QA | Quality Assurance |
| SCL | Spacecraft Command Language |
| SEL | Single Event Latch-up |
| SEU | Single Event Upset |
| SLAC | Stanford Linear Accelorator Center |
| TACK | Trigger Acknowledge |
| TDA | Tile Detector Assembly |
| T&DF | Trigger and Data Flow Subsystem (LAT) |
| TBD | To Be Determined |
| TBR | To Be Resolved |
| TSA | Tile Shell Assembly |
| TSS | Thermal Synthesizer System |
| TKR | The LAT Tracker Subsystem |
| VME | Versa Module Eurocard |
| WBS | Work Breakdown Structure |
| WOA | Work Order Authorization |

# Applicable Documents

Documents relevant to the ACD Photomultiplier Quality Plan include the following.

LAT-SS-00016, LAT ACD Subsystem Requirements – Level III Specification

LAT-SS-00352, LAT ACD Electronics Requirements – Level IV Specification

LAT-SS-00437, LAT ACD Mechanical Requirements – Level IV Specification

LAT-MD-00039-01, LAT Performance Assurance Implementation Plan (PAIP)

LAT-MD-00099-002, LAT EEE Parts Program Control Plan

LAT-SS-00107-1, LAT Mechanical Parts Plan

LAT-MD-00078-01, LAT System Safety Program Plan (SSPP)

ACD-QA-8001, ACD Quality Plan

LAT-TD-00760-D1 Selection of ACD Photomultiplier Tube

LAT-DS-00739-1 Specifications for ACD Photomultiplier Tubes

LAT-TD-00438-D2 LAT ACD Light Collection/Optical Performance Tests

LAT-TD-00720-D1 ACD Phototube Helium Sensitivity

LAT-DS-00740-1 Temperature Characteristics of ACD Photomultiplier Tubes

Response to RFQ 5-09742, Hamamatsu Photomultiplier Tube Proposal

- Introduction.

The idea of the test is to run the ACD with cosmic ray muons and obtain pulse-height histograms, corresponding to a MIP, for each tile.  These histograms will determine the tile light yield averaged over the tile area with muons distributed uniformly.  It was shown (LAT-TD-00438-D2, LAT ACD Light Collection/Optical Performance Tests) that the loss of 2-3 fibers will cause the shift of MIP pulse-height peak by 10-15% .  In testing the flight ACD, the readout will be gated by a VETO signal from ANY tile.  The task is to find which tile coincidence combination is best for analyzing each particular tile.  It can be done by simulating the ACD with the cosmic ray muon flux.  To prove this approach, both simulations and real measurements were performed with BFEM ACD.

- BFEM Muon test.

6 - upper

1 - lower

9             10

12               11

5 – upper                                            7 - upper

4 – lower                                            2 - lower

8 – upper

3 – lower

Fig.1 Tile Numbering for BFEM

The test was performed by pulse height analyzing the signals from each BFEM tile, gated by the signal from one of the tiles.  It was repeated twice, first gating with tile 11 (one of the top tiles), and second gating with tile 8 (an upper side tile).

Simulations were done by GEANT 3.21/FLUKA with the exact BFEM geometry.  The same tiles, 11 and 8, were used for event triggering.  The muon flux used for the simulations was taken from A. Stephens (fig.2).

Fig.2  Muon flux angular dependence (E(>200 MeV)

Comparison of simulation and test results.  Table 1 shows the fraction of triggers showing signals in each tile.  The pulse-height distributions obtained in the muon test and in the simulations, both triggered by tile 8, are shown in fig. 3 and fig.4, respectively.  (Similar results triggered by tile 11 are not shown here.)  Comparison between simulation and test results show that they are consistent, and that the simulations can be used to develop the muon test technique for the flight ACD configuration.

Table 1.  Fraction of triggers accompanied by a signal in each BFEM tile

| Tile number | Test – gated by 11 | Simulation – gated by 11 | Test – gated by 8 | Simulation – gated by 8 |
| --- | --- | --- | --- | --- |
| 1 | 0.04 | 0.043 | 0.049 | 0.056 |
| 2 | 0.055 | 0.068 | 0.049 | 0.061 |
| 3 | 0.067 | 0.079 |  |  |
| 4 | 0.049 | 0.042 | 0.052 | 0.062 |
| 5 | 0.065 ? | 0.018 | 0.068 | 0.083 |
| 6 | 0.022 | 0.018 | 0.028 | 0.023 |
| 7 | 0.09 | 0.17 | 0.066 | 0.084 |
| 8 | 0.21 | 0.17 |  |  |
| 9 |  |  | 0.018 | 0.019 |
| 10 |  |  | 0.017 | 0.018 |
| 11 |  |  | 0.11 | 0.18 |
| 12 |  |  | 0.11 | 0.17 |

Fig.4  Simulations: Histograms from BFEM tiles triggered by tile 8

Fig.3 Muon test: Histograms from BFEM tiles triggered by tile 8

- Test of the Flight ACD .  The task is to find for each particular tile what other tiles can be used to trigger its analysis.  The trajectories should be as normal to the tested tile surface as possible, with reasonable statistics to be collected from cosmic muons.  The tile numbering used in the simulations is shown in Fig. 5.  The simulation run corresponded to approximately 40 minutes of ACD running time, with 10-11 thousand triggers collected for each top tile.  For each tile, the triggering tiles were carefully selected, and corresponding histograms are shown in figures below.  For reliable fitting and MIP peak position determination, approximately 1,500 events are desirable in the histogram.  Looking at the histograms, we see that the most difficult tiles to calibrate will be the upper side tiles (fig. 7), which will require 6-7 hours to obtain ~1,500 events.  Limited calibration can be done within ~4 hours.

Tile T11 – middle of top edge                                       Tile T12 – next to the top central tile

Tile T13 – top center tile                                Tile T17 – diagonal from the top central tile

Fig.6 - Top tiles

Fig. 7 - Side tiles

Tile Z11 – end tile in the side top row                                Tile Z13 – middle tile in the side top row

Tile Z21 – end tile in side second row                  Tile Z22 – second tile from end in side second row

Tile 31 – end tile in side third row                         Tile Z32 – second tile from end in side third row

Tile Z41 - Long tile, side bottom row

Some remarks on the use of these histograms:

- Look at the quality of the histogram.  For those that have too few of events, project how would it look at higher statistics.

- These histograms would be collected in ~40 min of running time, so one can estimate how much time will be needed to get a given number of events.

Note which tiles were used for selecting events in each histogram.  Any suggestions to improve these choices would be welcome.

How many events do we need in the histogram for given peak position uncertainty? I believe that the common mathematical approach to this estimate is complicated by the high variability of the particle paths in the tile and the desired reduction of the number of events needed.  Simulations seem to be the appropriate way to do this analysis.  I did the following: Using the sea level muon flux in the simulations, the simulated pulse height distribution was fitted by a Landau distribution to find the peak position.  This was repeated for 10 sets of approximately 2,500, 1,000, and 500 events in the histogram and the mean value and standard deviation (() was determined for each of these sets of 10 runs. This was done for the tile on the top of ACD (for which most of muons hit the tile around normal incidence).  The results are given in Table 2.  The examples of the pulse height distribution for the top tile, with 996 events and 511 events in the histograms, are given in fig. 8 and fig.9 (column 3 and 4 in the table) respectively.  The histogram for the side (“bad”) tile (374 events) is shown in fig. 10 (column 5 in the table). It is seen that the precision of the peak position fitting is surprisingly high, even for such small statistics as ~500 events.  For the “bad” tile (the side one), there is a large variation of incident muon angles, and consequently the muon paths in the tile.  Note that we are looking for a change in the light yield of 5% and more.

Table 2   Simulations of the peak position determination precision

| Fitted MIP peak position | Top tile, ~2,500 evts | Top tile, ~1,000 evts | Top tile, ~500 evts | Side tile, ~400 evts, gain=1 | Side tile, ~400 evts, gain=0.95 |
| --- | --- | --- | --- | --- | --- |
|  | 330.8 | 329.9 | 332.2 | 487.3 | 467.4 |
|  | 333.0 | 326.8 | 338.2 | 495.2 | 467.2 |
|  | 336.7 | 336.6 | 324.5 | 510.3 | 462.8 |
|  | 340.9 | 338.4 | 331.5 | 481.3 | 489.9 |
|  | 330.7 | 335.0 | 340.9 | 482.1 | 478.9 |
|  | 334.9 | 330.9 | 348.4 | 491.9 | 539.5 ( |
|  | 331.0 | 336.0 | 326.8 | 509.5 | 447.1 |
|  | 336.8 | 339.1 | 343.0 | 510.9 | 482.1 |
|  | 331.2 | 339.5 | 336.5 | 511.7 | 482.5 |
|  | 336.2 | 334.8 | 350.8 | 521.1 | 476.4 |
| Mean (( | 334.2(3.4 (1%) | 334.7(4.3 (1.3%) | 337.3(8.7 (2.6%) | 500.1(13.8 (2.8%) | 478.7(24 (5%) |

Fig. 8   Top tile, 996 events       Fig.9  Top tile, 511 events     Fig. 10  side tile, 374 events

- The back-up option for testing the flight ACD. Another way to do the gain calibration test is to look at all tile histograms in a muon self-triggering mode, meaning that all signals recorded in given tile, will be used for the analysis.  The advantage of this approach is that for ~1 hour of instrument running there will be from 3,000 to 15,000 events in the histograms (depending on the tile), which provides a very reliable and precise peak position determination.  The disadvantage of this approach is that the histograms for some tiles, especially the side tiles, will be very dependent on the muon angular distribution.  No external triggering will be used, so the angular range of particles causing the triggering will be 2( for every tile.  The incident flux angular dependence could cause uncertainty if that angular distribution varies.  The muon flux angular distribution is constant for a given location, so this particular approach can be successfully used for repeated functional tests performed in the same place.  After moving to another place, re-calibration must be done using the approach described earlier, similar to fig. 6 and 7, with histograms selected by appropriate triggering.  The histograms for the same tiles, in self-triggering mode, are presented in fig. 11 and 12, which illustrate the high statistics achieved with this approach.

Tile T11 – middle tile in the edge row     Tile T12 – next to the central tile

Fig. 11 – Top tiles in self-triggering mode

Tile T13 – top center tile                                    Tile T17 – diagonal from the central tile

Tile Z11 – end tile in the top row        Tile Z13 – middle tile in the          Tile Z21 – end tile in the

top row                                           second row

Fig. 12 – Side tiles

How sensitive is the self-triggering mode to a gain change? A gain change of 5% was simulated.  A similar situation was experimentally tested and proven with muons on BFEM.  A simulation identical to that given above, but with a gain of 0.95 was performed.  The MIP peak positions for these runs are given in the Table 3 along with the statistics in each corresponding histogram.

| Tile | Statistics for 40 min | Peak position for gain = 1 | Peak position for gain = 0.95 | Ratio |
| --- | --- | --- | --- | --- |
| T11 | 10,650 | 498.9 | 472.7 | 0.947 |
| T12 | 10,680 | 502.2 | 476.4 | 0.949 |
| T13 | 10,828 | 497.7 | 471.6 | 0.947 |
| T17 | 10,651 | 497.4 | 471.6 | 0.948 |
| Z11 | 3,631 | 785.6 | 748.9 | 0.953 |
| Z13 | 3,601 | 793.0 | 756.7 | 0.954 |
| Z21 | 2,547 | 784.6 | 748.8 | 0.954 |
| Z22 | 2,512 | 776.2 | 729.7 | 0.940 |
| Z31 | 1,930 | 739.8 | 711.2 | 0.961 |
| Z32 | 1,944 | 765.5 | 735.8 | 0.961 |
| Z41 | 9,658 | 776.8 | 738.0 | 0.950 |

The results obtained demonstrate that the sensitivity of this approach is quite adequate (5% gain change detectable).

- Conclusion.

I believe that the gain calibration test should be done as follows:

- During ACD I&T –

a) gain calibration with muon hodoscope for each tile.  This will be the most precise measurement (almost free from the uncertainty introduced by the different muon arrival direction).  These results will serve as a reference in case of unclear future test results.

- a muon run for 8-10 hours, with triggering from any tile.  The results will be treated in two ways – selecting triggers, and self-triggering.  In both cases, the statistics reduction will be used to understand the stability of the results

2. In all other ACD test –Muon runs will be used for the available time, and depending on that time, the analysis approach will be selected.  If visible performance change is noted, more careful testing must be performed, possibly requiring more time.  In extreme cases, when tile replacement is contemplated, a muon hodoscope test should be performed before making the replacement decision.

# Text boxes

### Fig. 5

Tile Z22 – second tile from end in        Tile Z31 – end tile in 3-rd row   Tile Z41 – Long tile

side second row                                                                                        bottom row