|  | Document # | Date Effective |
| --- | --- | --- |
|  | LAT-TD-00843-D1 | Draft 07/17/02 |
|  | Prepared by(s) | Supersedes |
|  | Alex Moiseev | None |
| GLAST LAT Subsystem Technical Document |  |  |
|  | Subsystem/Office |
|  | Anticoincidence Detector Subsystem |
| Document Title |
| Design Qualification Tests for ACD TDA and Phototubes |

Gamma-ray Large Area Space Telescope (GLAST)

Large Area Telescope (LAT)

Design Qualification Tests for ACD TDA and Phototubes

DRAFT

CHANGE HISTORY LOG

| Revision | Effective Date | Description of Changes |
| --- | --- | --- |

# Purpose

This study reports on the performance of ACD tiles of different configurations, including the use of the fiber-to-fiber optical connectors, as well as some qualification tests of phototubes.

- Definitions and Acronyms

| ACD | The LAT Anti-Coincidence Detector Subsystem |
| --- | --- |
| ADC | Analog-to-Digital Converter |
| AEM | ACD Electronics Module |
| ASIC | Application Specific Integrated Circuits |
| BEA | Base Electronics Assembly |
| CAL | The LAT Calorimeter Subsystem |
| DAQ | Data Acquisition |
| EGSE | Electrical Ground Support Equipment |
| EMC | Electromagnetic Compatibility |
| EMI | Electromagnetic Interference |
| ESD | Electrostatic Discharge |
| FM | Flight Module |
| FMEA | Failure Mode Effect Analysis |
| FREE | Front End Electronics |
| GAFE | GLAST ACD Front End – Analog ASIC |
| GARC | GLAST ACD Readout Controller – Digital ASIC |
| GEVS | General Environmental Verification Specification |
| GLAST | Gamma-ray Large Area Space Telescope |
| HVBS | High Voltage Bias Supply |
| ICD | Interface Control Document |
| IDT | Instrument Development Team |
| I&T | Integration and Test |
| IRD | Interface Requirements Document |
| JSC | Johnson Space Center |
| LAT | Large Area Telescope |
| MGSE | Mechanical Ground Support Equipment |
| MLI | Multi-Layer Insulation |
| MPLS | Multi-purpose Lift Sling |
| PCB | Printed Circuit Board |
| PDR | Preliminary Design Review |
| PMT | Photomultiplier Tube |
| PVM | Performance Verification Matrix |
| QA | Quality Assurance |
| SCL | Spacecraft Command Language |
| SEL | Single Event Latch-up |
| SEU | Single Event Upset |
| SLAC | Stanford Linear Accelorator Center |
| TACK | Trigger Acknowledge |
| TDA | Tile Detector Assembly |
| T&DF | Trigger and Data Flow Subsystem (LAT) |
| TBD | To Be Determined |
| TBR | To Be Resolved |
| TSA | Tile Shell Assembly |
| TSS | Thermal Synthesizer System |
| TKR | The LAT Tracker Subsystem |
| VME | Versa Module Eurocard |
| WBS | Work Breakdown Structure |
| WOA | Work Order Authorization |

# Applicable Documents

Documents relevant to the ACD Photomultiplier Quality Plan include the following.

LAT-SS-00016, LAT ACD Subsystem Requirements – Level III Specification

LAT-SS-00352, LAT ACD Electronics Requirements – Level IV Specification

LAT-SS-00437, LAT ACD Mechanical Requirements – Level IV Specification

LAT-MD-00039-01, LAT Performance Assurance Implementation Plan (PAIP)

LAT-MD-00099-002, LAT EEE Parts Program Control Plan

LAT-SS-00107-1, LAT Mechanical Parts Plan

LAT-MD-00078-01, LAT System Safety Program Plan (SSPP)

ACD-QA-8001, ACD Quality Plan

LAT-TD-00760-D1 Selection of ACD Photomultiplier Tube

LAT-DS-00739-1 Specifications for ACD Photomultiplier Tubes

LAT-TD-00438-D2 LAT ACD Light Collection/Optical Performance Tests

LAT-TD-00720-D1 ACD Phototube Helium Sensitivity

LAT-DS-00740-1 Temperature Characteristics of ACD Photomultiplier Tubes

Response to RFQ 5-09742, Hamamatsu Photomultiplier Tube Proposal

- The experimental approach.

M1

S1

S2

T1

T2

S3

M2

Fig.1 Principal Scheme of the tests

Tile performance is measured by collecting the pulse-heights of signals created by cosmic muons in the tile.  For the tile efficiency measurement, the muon selection scheme has to be reliable, with the level of false triggers less than 1:10,000.  This was achieved by triggering the data acquisition by the coincidence of two triggering scintillators (M1 and M2), and selecting the events for the efficiency analysis by applying tight muon-like pulse-height requirements to the signals from S1, S2, and S3.  Signals from tiles T1 and T2 are used for the efficiency measurements.

- Earlier Tests, for Reference (from LAT-TD-00438-D2)

The absolute efficiency of a TDA (Tile detector assembly) has been measured, and the corresponding light yield estimated.  The measurements have been performed with relatively short fibers (only WSF), which collect the light from the tile and deliver it to the PMT.  In the real design there will be fiber bundles up to 1.5 meters long to deliver the light from the most distant tiles (in the middle of the top ACD surface) to the PMT location.  Two cases are considered: (1) the resulting efficiency if no clear fibers are used and light is delivered to a PMT by only WSF (1.5 meters long, 40% light loss assuming the published absorption length); and (2) the use of fiber-to-fiber connectors and clear fibers to deliver the light to PMT.   (Clear fibers are assumed to be much more transparent, with light loss assumed to be ~15% due to losses in the connector between the WSF and clear fibers.  NOTE: This assumption later proved to be incorrect; at least initially, losses in the clear fibers seem to be nearly as great as those in WSF.)

These tests were done using a Goddard-made tile and commercial Hamamatsu R647 phototubes.  The results, shown in Fig. 2 and Fig. 3, show the measured efficiencies for one and two phototubes on a single tile, along with the calculated light losses for longer fibers or connectors to clear fibers.  Light loss in the clear fibers was assumed to be nil

(but see the note just above).

- Test of FNAL first flight-prototype tile

The task was to see the performance of the first FNAL tile, compared to the tiles made at Goddard.  This tile actually has all our potential improvements, and is expected to have about 30% light yield increase compared with my flight prototypes (due to aluminization of the fiber ends).  The results for one of two PMTs (from one tile) running are shown in fig. 4, and for both PMT running in “OR” in fig. 5.  The black lines labeled by “FNAL XK 0515” and “FNAL XK 2028” in Fig. 4 correspond to two bundles from the first FNAL tile, viewed by R647 commercial PMT.  In Fig. 5, the red line corresponds to an OR of the signals from these two tubes.  The pink dashed lines is a fitted efficiency, from a Poisson distribution assuming some number of photoelectrons (p.e).  The superior performance of the new tile made at Fermilab is demonstrated by comparison to the black lines in Fig. 2 and 3 above, made with the GSFC tile.  These measurements were all made close to the tile, with no long fibers.

- Test of the First Qualification PMTs.

Figures 4 and 5 also show results for the first qualification-type R4443 phototubes from Hamamatsu.  In Fig. 4, the lines labeled by 887 (green), 889 (black), and 882 (red), show single-phototube efficiency measurements using the GSFC tile and the R4443 tubes with these serial numbers.  These curves can be compared to the black curve in Fig. 2, made using a commercial tube.   As can be seen, the 889 and 887 tubes do not perform as well as the commercial tube.  In consultation with Hamamatsu, we learned that these tubes have a lower quantum efficiency than the 882 tube that does produce results as good as the commercial tube.  For this reason, we have added to our phototube specification the requirement that the quantum efficiency at 490 nm (the waveshifting fiber peak output) be no less than 15.5%.  Using this criterion, tubes 889 and 887 were rejected, while tube 882 was accepted.  The OR of tubes 882 and 887 attached to the GSFC tile is shown as the black curve in Fig. 5.  Despite the lower performance of the 887 tube, the combination of the two produces an efficiency that is slightly higher than that measured with the commercial tubes (the black curve in Fig. 3).  These tests confirm that the R4443 tubes meet the ACD efficiency requirements.

- Tests of different tiles and fiber coupling to phototubes.

In these tests the position of the MIP peak position in the pulse-height histogram was determined, using the same scheme of the measurements as described above, for a variety of tiles and light collection methods. The measurement results are summarized in Table 1.  Three important conclusions can be drawn from this table:

- The bent tile performs as well as the flat tile, even through the bent region.  This test confirms that the bending does not cause an appreciable loss of light.

- The light loss through the optical fiber-to-fiber connector and a 1.2 m long clear fiber bundle is substantial.  It was very discouraging to see such a big light loss in case of green fiber coupling to the clear fibers (~40% light yield reduction).  This was true whether the polishing of the fiber ends was done by hand or with a diamond mill.  The loss is far greater than that estimated in section 2 above, which assumed the principal loss to be 15% loss in the optical connector.

- Thermal bonding of the waveshifting fibers to clear fibers produces no better results than connecting the fibers through the optical connector.  The problem seems to be independent of the coupling method.

Table 1.  Pulse Heights (channels) for various combinations of tile, fibers, and PMTs

| MIP peak position | FNAL first tile (for this tile the efficiency given above was measured) | FNAL bent tile, center of the tile | FNAL bent tile, area of bending | FNAL tile with fiber-to-fiber connector to clear fibers, hand fiber ends polishing | FNAL tile with fiber-to-fiber connector to clear fibers, diamond milled fiber ends | FNAL tile with thermally coupled clear fibers |
| --- | --- | --- | --- | --- | --- | --- |
| PMT XK 0515 | 469.3 (measured 3 times) | 486.9 | 497.5 | 243.5 | F 254.5 | 221.1 |
| PMT XK 2028 | 527.3 (measured 3 times) | 542.8 | 552.8 | 267.7 | 281.4 336.0 331.0 571.0 | 232.8 309.0 328.0 496.9 640.0 |

A long-term (4 days) efficiency measurement was performed on the tile with the optical connector and clear fiber bundle to be sure that the efficiency degradation is consistent with the MIP peak shift.  This test represents a complete end-to-end test of a flight-like Tile Detector Assembly:  scintillator, waveshifting fiber, optical connector, clear fiber, photomultiplier tube.  The results are shown in fig. 6.  For the individual tubes, the measured efficiency at the nominal threshold of 0.3 MIP is above 0.9990 but below the ACD efficiency requirement of 0.9997.  For the combination of the two tubes, the efficiency meets the ACD requirement with some margin.  From a comparison with Figures 2 and 3 it can be seen that the loss of light in the connector and clear fibers offsets the gains made with the improved Fermilab tile.

In Table 1, comparison should be made between lines measured with first FNAL tile with no optical connectors, shown on figures 4 and 5 (“FNAL XK 0515” and “FNAL 2028”) and lines labeled “XK 0515” and “XK 2028” on fig. 6 measured for the tile with optical connector and clear fibers.  Comparison of estimated light yield (26-30 p.e. for the former assembly, and 15-18 p.e. for the latter one) demonstrates a good agreement with muon peak shift.  This comparison shows the validity of measuring the muon peak as a tracer of performance rather than always requiring a complete efficiency measurement, which typically requires much longer observation.

- Tests of absorption in clear fibers.

In attempting to trace the loss of light in the full optical assembly, we measured the peak position of the MIP peak in scintillators using varying lengths of clear fiber.  The results are summarized in Table 2.  It seems obvious from these measurements that the clear fibers we are using (Bicron, nominally multiclad 1.2mm, but Bicron now claims they have never made such fibers.  They may be single-clad instead of multiclad.), have a significant absorption, probably more than implied by the nominal 12 m absorption length quoted by Bicron.  It appears that at least as much light is lost in the 1.2 m fibers as in the optical connector.

Table 2.  Pulse Heights (channels) with different length of clear fibers, compared with clear fibers and connector removed:

|  | FNAL tile with connector | FNAL tile with thermal splices |
| --- | --- | --- |
| PMT XK 0515 FNAL polishing 1.2m clear fibers B.D. polishing, 1.2m clear fibers B.D. polishing, 20cm clear fibers B.D. polishing, no connector | 304 296 - 530 | 278 428 559 |
| PMT XK 2028 FNAL polishing, 1.2m clear fibers B.D. polishing, 1.2m clear fibers B.D. polishing, 20cm clear fibers B.D. polishing, no connector | 336 331 571 | 309 328 497 640 |

A 1.8 m sample of the clear fiber was sent to Fermilab for an independent measurement of the absorption.  The results are shown in Fig. 7

Fig. 7  Light absorption in clear and waveshifting fibers (relative).  The three points for the clear fibers suggest an absorption length (1/e length) of about 5 m.  An 18% light loss is seen between 0.9 m length fibers and 1.8 m length fibers.

Although the absorption is significantly larger in the waveshifting fibers, the additional loss in the optical connector implies that a careful tradeoff must be done between longer waveshifting fibers and an optical connector/clear fiber approach.  A preliminary analysis suggests that the break-even point comes at about 0.7 m.  Any tile closer than this to its phototubes would benefit from using waveshifting fibers only.

A concern is that the absorption in the clear fibers is much larger than advertised by Bicron.  There is some possibility that the clear fibers were single clad rather than multi-clad (Bicron says they have no record of selling us multi-clad fibers).  Additional clear fibers have been ordered, explicitly multiclad.  Further testing is needed before the design is finalized.

- Conclusion

Figure 6 gives us an understanding of the real light budget.  It appears that we lost all our gains for last 1.5 years (we are back to 15-18 p.e.), but we are still meeting our requirements (but with little margin).  If we look at the fig.7, where the light attenuation for green fibers (WSF) is shown (preliminary measurements), we see that only for the central top tiles will the light loss due to the connector be comparable to that of the attenuation in the green fibers.  For all other tiles, the use of only green fibers looks preferable; however, the fabrication, shipping and assembly will be a nightmare.  The light loss in the clear fibers should be studied, as well as the reason our connectors give such a big loss (FNAL data give 12-15% loss in similar connectors).

# Text boxes

Fig. 6.  Efficiency measurement for a complete flight-like assembly.  The curves to the left are for individual phototubes connected to one flight-type scintillator tile through waveshifting fibers, an optical connector, and a clear fiber bundle.  The curve to the right is the OR of the signals from the two tubes.

Fig.4. Light budget for 1 PMT.

FNAL lines – measured efficiency using commercial phototubes on the FNAL tile, with corresponding p.e. calculations.

Other lines – measured efficiency using various phototubes on the GSFC tile, with corresponding p.e. calculations.  These curves compare to the black line in Fig. 2.

Fig.2. Light budget for 1 PMT.

Black line – measured efficiency (corresponds to 19 p.e.)

Green line – efficiency assuming 15% light loss (16 p.e.) in the fiber connector and clear fibers

Blue line – efficiency assuming 40% light loss (12 p.e.) in 1.5 meters WSF

Fig.3. Light budget for 2 PMTs.

Black line – measured efficiency

Green line – efficiency assuming 15% light loss

Blue line – efficiency assuming 40% light loss

Fig.5. Light budget for 2 PMTs.

FNAL line – measured efficiency using 2 commercial phototubes on the FNAL tile, with corresponding p.e. calculations.

Black line – measured efficiency using 2 phototubes on the GSFC tile, with corresponding p.e. calculations.  This curve compares to the black line in Fig. 3.