LCLS

Phase and Amplitude Controller

PAC

Jeff Olsen

03/17/06

2K samples X 16 bits = 32K bits * 2 = 72Kbits of memory

360 Hz = 2.77mS

Switching to uCDIMM  processor board

3 Trigger Modes

00 – External Trigger

01 – Internal Trigger

10 – CPU Trigger

Voltages to monitor

|  | Source | Type | Range |  |
| --- | --- | --- | --- | --- |
| 0 | SSSB Amp | Temperature |  |  |
| 1 | IQ Modulator | Temperature |  |  |
| 2 | Control Board | Temperature |  |  |
| 3 | SSSB Amp | Bulk Supply | +12V |  |
| 4 | SSSB Amp | Bulk Supply | -12V |  |
| 5 | RF Board | Bulk Supply | +15V |  |
| 6 | Control Board | Bulk Supply | +5V |  |

ColdFire Configuration

The PAC board interface to the Coldfire uses Chip Select, CS1 and Byte Select  BS3 and BS2. To properly configure the Coldfire, Port J Register, (PJPAR)  must be configured for Byte Select and Chip Select functionality, not IO. This is done by writing an 0xFF to address 0x40100054. Then you must configure the Chip Select Module (CSM) for proper decoding of the base address.

ColdFire Configuration

|  | Register | Value | Comment |
| --- | --- | --- | --- |
| 0x40100054 | PJPAR | 0xFF | Configure PJPAR for Byte and Chip Select, NOT IO |
| 0x4000008C | CSAR1 | 0x80000000 | PAC Base Address |
| 0x40000090 | CSMR1 | 0x000000001 | Sets Valid bit |
| 0x40000096 | CSCR1 | 0x000000A0 | Set for 16 bit, Not internal TA, Enable BS, No Burst. |

Memory Map

| Register Number | Memory Address | Use |  |  |
| --- | --- | --- | --- | --- |
|  | 00000 0FFE | I Memory | 16 bit Signed Integer |  |
|  | 1000 1FFE | Q Memory | 16 bit Signed Integer |  |
| 0 | 2000 | Control Register | 16 bit unsigned |  |
| 1 | 2002 | DAC Delay | 16 bit unsigned |  |
| 2 | 2004 | SSSB Delay | 16 bit unsigned |  |
| 3 | 2006 | SSSB Width | 16 bit unsigned |  |
| 4 | 2008 | Ext Trigger Delay | 16 bit unsigned |  |
| 5 | 200A | Ext Trigger Width | 16 bit unsigned |  |
| 6 | 200C | Int Trigger Rate | 16 bit unsigned |  |
| 7 | 200E | Clock Select | 16 bit unsigned |  |
| 8 | 2010 | PAC Temperature | 16 bit unsigned |  |
| 9 | 2012 | IQ Temperature | 16 bit unsigned |  |
| 10 | 2014 | SSB Temperature | 16 bit unsigned |  |
| 11 | 2016 | VFWD | 16 bit unsigned |  |
| 12 | 2018 | VREFL | 16 bit unsigned |  |
| 13 | 201A | SSB_PWR | 16 bit unsigned |  |
| 14 | 201C | +5V | 16 bit unsigned |  |
| 15 | 201E | -12V | 16 bit unsigned |  |
| 16 | 2020 | CPUTrig address |  |  |

Temperature Conversion:

ADC input is 0 to 2.5V, Temp input is 20C to 70C, ADC output is 0 to 16383 counts so the conversion is

T (C) = (Count / 16384) * 2.5 * 20 + 20

Voltage Conversion:

ADC input is 0 to 2.5V, Input voltage is 0V to 20V so the conversion is

V(in) = (Count / 16384) * 20

The -12V ADC reading will show a positive 12V since the input is inverted.

Control Register (2000)

| Bit |  |  | R/W |  |
| --- | --- | --- | --- | --- |
| [15..08] | Version Number |  | R |  |
| [07..06] | Trigger Source | ‘11’=> StandBy Trigger ‘10’ => CPU Trigger ‘01’ => Internal Trigger ‘00’ => External Trigger | R |  |
| 05 | Stby Trigger Missing | No transition on Stby Trig in 1sec Latched value Writing a ‘1’ => Clears flag | R/W |  |
| 04 | 119Mhz Missing | Check 119Mhz clock for 90-150Mhz. Latched value Writing a ‘1’ => Clears flag | R/W |  |
| 03 | Ext Trig Missing | No transition on Ext Trig in 1sec Latched value Writing a ‘1’ => Clears flag | R/W |  |
| 02 | SSBOverTemp | SSB Over Temperature Latched value Writing a ‘1’ => Clears flag | R/W |  |
| [01..00] | Trigger Mode | ‘10’=> CPU Trigger ‘01’ => Internal Trigger ‘00’ => Normal; External Trigger | R/W |  |

Clock Select (2007)

| Bit |  |  | R/W | Default |
| --- | --- | --- | --- | --- |
| [15..01] | Unused |  | R |  |
| [00] | Clock Select | ‘1’ => Internal CPU CLK *2 ‘0’ => External 102Mhz |  | ‘0’ |

LED’s

| D1 | 119Mhz missing | Current status of 119Missing OneShot |  |  |
| --- | --- | --- | --- | --- |
| D2 | Ext Trig missing | Current status of EXT Trig Missing OneShot |  |  |

Connector Assignments

| H1 | SSSB OverTemp |  | Input | 0-20V |
| --- | --- | --- | --- | --- |
| H7 | Trigger Out |  | Output | TTL |
| H9 | I Monitor |  | Analog |  |
| H10 | Q Monitor |  | Analog |  |
| H11 | Reset |  |  |  |
| H12 | IQ Modulator Temperature | AD592 | Temp |  |
| J1 | Ethernet Port |  |  |  |
| J2 | 119Mhz | 10Dbm (min) | Input | RF |
| J3 | External Trigger |  | Input | NIM |
| J4 | Q Out |  | Analog | 50 ohm |
| J5 | I Out |  | Analog | 50 ohm |
| J6 | SSB Gate | Copy of P5-12 |  |  |
| J7 | Trigger Out | Copy of H7 | Output | TTL |
| J11 | Alternate Ethernet |  |  |  |
| J10 | Alternate Trigger |  |  |  |
| P1 | RS232 Port 0 | From Arcturus |  |  |
| P5 | SSB IO |  |  |  |
| P6 | Power Connector |  |  |  |

P5 Connector Pinout

| Pin |  |  |  |  |
| --- | --- | --- | --- | --- |
| 1 | SSB Temperature | AD592 | Temp | I out |
| 2 | SSB Temperature |  | Temp | Return |
| 6 | SSB VFWD |  | Analog |  |
| 8 | SSB VREFL |  | Analog |  |
| 10 | SSB V+ |  | Analog |  |
| 12 | SSB Gate | Copy of H6 | Output | TTL |
| 14(*) | SSB OverTemp | Same as H1 | Input | 0-20V |

(*) Wire modification on Rev 1 board

P6 Connector Pinout

| Pin |  |  |  |  |
| --- | --- | --- | --- | --- |
| 1 | +5V |  |  |  |
| 2 | +5V |  |  |  |
| 3 | +5V |  |  |  |
| 4 | -12V |  |  |  |
| 5 | GND |  |  |  |
| 6 | GND |  |  |  |
| 7 | GND |  |  |  |
| 8 | +12V |  |  |  |

Xilinx Version History

|  | Version |  |  |  |
| --- | --- | --- | --- | --- |
| 11/08/06 | 04 | Recompiled firware under ISE8.2 |  |  |
|  |  | Having problems getting over 110Mhz with new ISE??? ReWrote prog_strb using 1 up counter. Got 133Mhz |  |  |
|  | 04inv | Discovered glitch on the Q output. Inverted Clock to see if this fixes the timing problem with the DAC clock | OK |  |
| 11/08/06 | 05 | Added logic for Stby Trigger |  | Compile Time ~4min |
| 03/13/07 | 06 | Added clock mux for incoming RF and CPU clk |  |  |
|  |  | Added registers to control this bit. (2007) |  |  |
|  |  | Fixed error in latching the Trigger Mode status |  |  |
| 03/21/07 | 07 | Fixed error in Ext/Stby trigger logic |  |  |
|  |  | Changed ClkSelect to default to External Clk |  |  |