Temperature Stability of RF Components

R. Akre 4/21/2005

Typically the phase change of the RF passing through a device will change with temperature.  With near speed of light cables, the cable length increases with temperature and the phase goes negative.  Changes in dielectric constant with temperature will also effect the phase changes.  The specification for LCLS to achieve 70fS rms error between the phase of the RF and the electron beam is not easily realizable.  Some characteristics of the task and the technical approach are described in previous papers, see references.  Measurements of the RF system have shown portions of the system to be stable to LCLS specifications for periods of 15 minutes.  Temperature changes will drive the RF system out of tolerance beyond that.  The phase of the RF which the beam sees in an accelerator structure can't be measured with a single RF measurement.  The phase change through the structure is too large to measure to within LCLS tolerances.  We will measure the input phase, the output phase and the temperature of the structure along with beam based parameters, energy and bunch length, to determine the phase and amplitude which the beam saw passing through the accelerator.

We have received new low noise oscillators of which one will be installed to run the 2 mile linac.  One of the other oscillators will be used to measure phase noise of the components in the 2 mile linac over the next few months.  This will determine which components, in the phase reference system, will need to be upgraded.  The phase noise of these oscillators is shown below.

Below is a portion of the data collected over the past few years on linac RF phase stability.  The amplitude of the RF pulse specifications are easier to achieve than the phase, so the efforts are concentrated on phase stability.  Changes in the amplitude will cause amplitude to phase conversion and the phase will go out of tolerance before the amplitude.

In the below data 1ºS (1º at 2856MHz) = 1pS or 1000fS.

A look at temperatures in the SLAC Linac

Temperature Variations

The temperatures of concern 		RF devices			1 Week STD	4hours

Linac Tunnel Temperature		RF cables				0.17F	.084F

Gallery Temperature			RF Cables				5.1F	0.9F

Accelerator Water (ACC_H2O)	Accelerator Structures			0.11F	0.052F

(See Appendix A also)		SLED Cavities

Waveguide Water (WG_H2O)	Low Level RF (Multipliers)		0.07F	0.045F

The following are plots of temperature variations over one week.

Linac Tunnel Temperature			Linac Gallery Temperature

Accelerator Water Temperatue		Waveguide Water Temperature

Appendix A shows the water temperature variations over 3 different time intervals.  The standard deviations are 0.059F over 14 minutes, 0.067F over 133 minutes, and 0.075F over 1 week.  These data were taken while the accelerator was in normal operation.  Sector 0 has the best temperature regulation and is the most critical sector in the accelerator.  This is the result of a greater effort being spent in tuning of the water temperature feedback loops in sector 0.  LCLS operation will request the same level of regulation in the sector 20 off axis injector.

RF Components, Their Temperature Coefficients and Environment

Accelerator Structures 			-8.6 ºS/ºF		ACC_H2O

SLED Cavities				+8.5 ºS/ºF		ACC_H2O

1-5/8 Rigid Coax (100 meters)		0.2 ºS/ºF		Tunnel

1/2 inch Heliax (100 meters)		2.9 ºS/ºF		Tunnel

WR284 Waveguide (100 meters)	4.6 ºS/ºF		WG_H2O

Frequency Multipliers			1.8 ºS/ºF		WG_H2O

PAD (Phase and Amplitude Detector)	0.02 ºS/ºF		WG_H2O

LCLS Phase Errors due to Temperature Changes over 4 hours

100 ft heliax in the tunnel		1 ºS/ºF x 0.084ºF = 0.084ºS

Accelerator Structures 10ft		-8.6 ºS/ºF x 	0.052ºF = 0.45ºS

SLED Cavities				+8.5 ºS/ºF x 0.052ºF = 0.44ºS

1-5/8 Rigid Coax (100 meters)		0.2 ºS/ºF x 0.084ºF	= 0.017ºS

WR284 Waveguide (100 meters)	4.6 ºS/ºF x 0.045ºF	 = 0.21ºS

Frequency Multipliers			1.8 ºS/ºF x 0.045ºF = 0.081ºS

PAD (Phase and Amplitude Detector)	0.02 ºS/ºF x 0.045ºF = 0.0009ºS

Most of the above temperature drifts will bring the LCLS RF system out of tolerances.  Feedback must be used to control the system.  Parts of the system have been shown to be stable for up to 15minutes without feedback.

References:

R. Akre, V. Bharadwaj, P. Emma, P. Krejcik, “SLAC Linac RF Performance for LCLS”, SLAC-PUB-8574, and THC11, LINAC 2000, Monterey, CA, Aug. 2000

R. Akre, P. Emma, P. Krejcik, " Measurements on SLAC Linac RF System for LCLS Operation", SLAC-PUB-8820, and TPAH105, PAC 2001, Portland, OR.

APPENDIX A

Accelerator Water Temperatures SLAC Linac

The 10 foot S-band accelerator structure has a phase slip of 15.5(S/(C(1) or 8.6(S/(F.  The temperature induced phase error of the RF as measured at the input of the structure or the output of the structure relative to the average phase the beam sees as it traverses the structure is ½ the above value, 4.3(S/(F.  If we assume the accelerator temperature variation is the same as the accelerator water temperature variation, in order to hold the beam to RF phase to within 0.1(S the temperature variation must be no greater than 0.023(F.  Sector 0, which is the most critical sector in the accelerator, would meet this requirement.

Accelerator Water Temperatures and Phase Errors due to Accelerator Structure Only

512 point data were taken with correlation plots.  During the second data set, 8000 seconds or 2.2 hours, sector 21 gallery temperature increased 4.3(F.

| Accelerator Water Temp |  |  | Deg F |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  | 512 points in 840Sec. | 512 points in 8000Sec | History Buffer | STD Phase |
|  | 3/6/99 6:25 |  | 3/6/99 7:00 |  | 1 week - end 3/6/00 | Error |
| Sector | AVG | STD | AVG | STD | AVG | STD | Deg S |
| 0 | 109.5 | 0.020 | 109.5 | 0.024 | 109.5 | 0.043 | 0.19 |
| 1 | 110.6 | 0.036 | 110.7 | 0.050 | 110.7 | 0.057 | 0.25 |
| 2 | 110.7 | 0.068 | 110.8 | 0.040 | 110.7 | 0.047 | 0.20 |
| 3 | 111.4 | 0.034 | 111.5 | 0.053 | 111.4 | 0.064 | 0.28 |
| 4 | 112.1 | 0.042 | 112.1 | 0.048 | 112.1 | 0.059 | 0.25 |
| 5 | 112.1 | 0.047 | 112.1 | 0.069 | 112.1 | 0.062 | 0.27 |
| 6 | 112.0 | 0.052 | 112.1 | 0.066 | 112.0 | 0.073 | 0.31 |
| 7 | 112.0 | 0.107 | 112.0 | 0.106 | 111.9 | 0.122 | 0.53 |
| 8 | 112.0 | 0.084 | 112.1 | 0.095 | 112.0 | 0.106 | 0.46 |
| 9 | 112.2 | 0.043 | 112.3 | 0.066 | 112.2 | 0.071 | 0.31 |
| 10 | 111.8 | 0.050 | 111.8 | 0.098 | 111.8 | 0.104 | 0.45 |
| 11 | 111.9 | 0.032 | 111.9 | 0.032 | 111.9 | 0.043 | 0.19 |
| 12 | 112.1 | 0.045 | 112.1 | 0.049 | 112.1 | 0.063 | 0.27 |
| 13 | 111.9 | 0.033 | 111.9 | 0.054 | 111.8 | 0.061 | 0.26 |
| 14 | 111.7 | 0.065 | 111.7 | 0.066 | 111.7 | 0.072 | 0.31 |
| 15 | 111.8 | 0.039 | 111.9 | 0.066 | 111.9 | 0.065 | 0.28 |
| 16 | 111.9 | 0.045 | 111.9 | 0.058 | 111.9 | 0.067 | 0.29 |
| 17 | 111.9 | 0.056 | 111.9 | 0.074 | 111.9 | 0.070 | 0.30 |
| 18 | 111.9 | 0.065 | 112.0 | 0.063 | 111.9 | 0.081 | 0.35 |
| 19 | 111.9 | 0.083 | 111.9 | 0.106 | 111.9 | 0.101 | 0.43 |
| 20 | 112.2 | 0.073 | 112.1 | 0.076 | 112.1 | 0.066 | 0.28 |
| 21 | 112.1 | 0.180 | 112.2 | 0.102 | 112.1 | 0.146 | 0.63 |
| 22 | 111.9 | 0.054 | 112.0 | 0.074 | 111.9 | 0.087 | 0.37 |
| Avg | 111.7 | 0.059 | 111.8 | 0.067 | 111.7 | 0.075 | 0.324 |

Teflon thermal expansion:    http://www.andersonmaterials.com/tma.htm

#### Thermomechanical Analysis (TMA) 
Thermal Analysis

The TMA uses interchangeable probes at varied loads to make a number of measurements, including the softening temperature or glass transition temperature, tensile modulus, compression modulus, coefficient of thermal expansion (CTE), melting temperature, crystalline phase transition temperature, and crystalline to amorphous transition temperatures. DSC can also be used to measure glass transition temperatures, crystalline phase transition temperatures, crystalline phase transformation temperatures, and melting point temperatures. Samples in the form of plugs, films, powders, or fibers, are compressed or held in tension by a probe assembly. Movement of the probe is translated electrically by a linear variable differential transformer (LVDT), attached to the probe, into a signal. Capable of testing polymers, metals, glasses, ceramics, inorganics, composites, and semiconductors. Temperatures can range from -180°C to 800°C. Sample dimensions: a maximum of 25 mm in height and 9.5 mm in diameter; a minimum of 0.1µm thick and 1.5 mm in diameter, an ideal thickness being 1-4mm. The applied load can range from 0-100 g. The change of dimension should be greater than 2x10-3µm/°C

ASTM Standard Test Methods: ASTM E228 - Standard Test Method for Linear Thermal Expansion of Solid Materials with a Vitreous Silica Dilatometer ASTM E831 - Standard Test Method for Linear Thermal Expansion of Solid Materials by TMA ASTM E1545 - Standard Test Method for Assignment of the Glass Transition Temperature by TMA

Illustrative Example: Rapid thermal expansion of PTFE crystalline phase TMA was used to characterize the properties of various PTFE powder materials. These were examined after pressing them into sheets. The sheet thickness was measured from -100°C to 200°C using a ramp rate of 4°C/min. The first of several TEFLON® materials (Fig. 1) exhibits an extensive crystalline transition at approximately 19°C due to a transformation of the Phase II helix crystal structure first to the Phase IV hexagonal crystal structure and then to the Phase I hexagonal crystal structure. This large peak created at transition indicates a PTFE which has a substantial crystalline phase and exhibits a drastic dimension change as it changes the crystalline phase. Also shown by the TMA is the softening point (Tg or Ts) of the TEFLON® at approximately 160°C. Data for another of several samples is seen in Fig. 2. This TEFLON® material exhibits a significantly smaller crystalline to amorphous transition dimension increase compared to the material observed in Fig. 1. This material initially consisted of a greater concentration of amorphous material. For our customer’s intended use, a rapid thermal expansion of a PTFE material at the crystalline transition temperature could not be tolerated. Their requirement was therefore for an amorphous PTFE material.

Thermoplastics - An Introduction

Topics Covered

Background Thermoplastic Characteristics Amorphous Polymers Semi-Crystalline Polymers

Background

What is an engineering thermoplastic? Such a definition is difficult to arrive at and is very subjective, but, for the purposes of this review article, any thermoplastic that can be formed into a load-bearing shape that might otherwise be formed from, for example, steel or wood will be classed as an engineering thermoplastic. This review is restricted to unreinforced and short fibre reinforced thermoplastics where the reinforcing fibre (usually glass or carbon) is typically less than 2-3 mm in length. Thermoplastics offer many advantages over traditional materials, including: low density; low energy for manufacture; low processing costs; and the ability to make complex shapes relatively easily. Thermoplastic Characteristics Thermoplastic materials generally fall within two classes of molecular arrangement, amorphous and semi-crystalline (see figure 1). Figure Error! Bookmark not defined. Molecular Arrangement of Polymer Chains Table 1 lists a selection of amorphous and semi-crystalline polymers. Table 1. Amorphous and semi-crystalline polymers. Amorphous Semi-crystalline Polyamideimide Polyetheretherketone Polyethersulphone Polytetrafluoroethylene Polyetherimide Polyamide 6,6 Polyarylate Polyamide 11 Polysulphone Polyphenylene sulphide Polyamide (amorphous) Polyethylene terephthalate Polymethylmethacrylate Polyoxymethylene Polyvinylchloride Polypropylene Acrylonitrile butadiene styrene High Density Polyethylene Polystyrene Low Density Polyethylene Amorphous Polymers Generally, fully amorphous polymers are stiff, brittle and clear in the virgin state. The temperature and stress state have a profound effect on the molecular arrangement and hence the properties of a polymer. Under the action of sufficient stress, the polymer chains can uncoil and align over a period of time. At elevated temperatures, polymer chains have enough energy to rotate and coil up further. Figure 2 shows schematically the effect of temperature on the elastic modulus of an amorphous and a semi-crystalline polymer. Below a temperature known as the glass transition temperature, Tg, the structure of amorphous polymers is termed 'glassy', with a random arrangement of the polymer chains, similar to the random molecular arrangement found in glass. As the temperature increases to Tg, the polymer chains have sufficient thermal energy to rotate, resulting in a drop in modulus. One definition of Tg is the temperature at which molecular rotation about single bonds becomes favourable as the temperature increases. At temperatures above Tg but below the melting temperature, Tm, there is a 'rubbery region', where the material can exhibit large elongations under relatively low load. Amorphous thermoplastics are generally used at temperatures below their Tg, where they can be brittle, just like glass. There are, however, certain exceptions. Polycarbonate (PC) is amorphous yet it is considered tough at temperatures well below its Tg. When it does fracture below its Tg, it does so in a brittle manner, but this requires a large amount of energy and so PC is considered tough, finding use in applications requiring impact resistance, e.g. safety helmets and bullet proof glazing. This behaviour is due to the chemical bonds in polycarbonate rather than the arrangement of the polymer chains. Figure 2. The effect of temperature on the Elastic Modulus (assuming both types of polymer have the same Tg and Tm) Semi-Crystalline Polymers Semi-crystalline materials such as polyamides do not exhibit a clear Tg or 'rubbery' region, although one is often quoted as the amorphous parts of the structure will undergo some transition. For these polymers the main transition occurs at Tm when the crystalline regions break down (see Fig. 2). Some chain rotation in the amorphous regions will occur below Tm, giving some impact resistance at these temperatures. Values of Tg and Tm for a number of polymers are given in Table 2. Table 2. Glass transition and melting temperatures for a range of thermoplastics. Polymer Tg (°C) Tm (°C) High density polyethylene (HDPE) - 135 Polypropylene (PP) -10 175 Polystyrene (PS) 100 - Polymethylmethacrylate (PMMA) 105 - Polyvinylchloride (PVC) 65 - Natural rubber (NR) -75 25 Polydimethyl siloxane (PDMS) -125 -55 Polyoxymethylene (POM) 125 175 Polycarbonate (PC) 150 - Polyethylene terephthalate (PET) 70 265 Polyetheretherketone (PEEK) 145 335 Nylon 6 (PA6) 50 215 Polyamideimide (PAI) 295 - Polysulphone (PSul) 195 - Polyphenylene sulphide (PPS) 90 285 Polyethersulphone (PES) 230 - Polyetherimide (PEI) 218 - Polytetrafluoroethylene (PTFE) 20 325 Liquid crystal polymer (LCP) - 420 Table 2. Values of Tg and Tm for selected polymers

Abstracted From the Materials Information Service, edited by Justin Furness For more information on Materials Information Service please visit The Institute of Materials